## Iodate dispenser
Calibration of Wheaton 10 ml bottle top dispenser used for dispensing standard solution.
Calibration preformed on Dec. 19, 2000 at CSIRO in Hobart
Dispenser was filled with DIW and 10 ml aliquots were dispensed into a small plastic beaker and weighed.
|  |  | rw |  | fbuoy
Temperature of DIW |  | 24 | 0.9973 |  | 1.001
Wt, g
9.9679
9.9638
9.9656
9.9682
9.9708
9.9687
9.9684
9.9684
9.9648
9.9685
|  |  | Vol 24C | Vol 20C
Avg | 9.9675 |  | 10.005 | 10.0046
SD | 0.0020047

## Calculations
Oxygen concentration calculations based on WHP Operations and Methods publication of 7/91
written by C. H. Culberson
O2 in ml/L=(((Vx-Vb)*VIO3*NIO3*5598)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
O2 in umol/L=(((Vx-Vb)*VIO3*NIO3*25000)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
Vx=endpoint of sample titration in uliters
Vb=reagent blank in uliters
VIO3=volume of IO3 standard in ml
NIO3=normality of iodate standard in equiv/L
Vstd=endpoint of the standard in uliters
DOreg=0.0017 ml O2 or 0.076 umoles O2 in volume of reagents added to samples
Vbot=volume of bottle in ml
Vreg=volume of reagents added (2 ml)

## O2
CTD/Rosette# | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L
7/23 | 156P | 2000-12-27 |  | 1391.23 | 643.6 | -0.1 | 141.7179 | 386.6849 | 8.6584
7/21 | 110P | 2000-12-27 |  | 1305.73 | 643.6 | -0.1 | 139.9933 | 367.4242 | 8.2271
7/20 | 132P | 2000-12-27 |  | 1308.27 | 643.6 | -0.1 | 142.2252 | 362.2805 | 8.112
7/19 | 131P | 2000-12-27 |  | 1300.94 | 643.6 | -0.1 | 143.7849 | 356.285 | 7.9777
7/18 | 103P | 2000-12-27 |  | 1239.7 | 643.6 | -0.1 | 142.5028 | 342.5872 | 7.671
7/17 | 124P | 2000-12-27 |  | 1219.34 | 643.6 | -0.1 | 150.2552 | 319.3329 | 7.1503
7/15 | 114P | 2000-12-27 |  | 1051.3 | 643.6 | -0.1 | 140.4964 | 294.6532 | 6.5977
7/14 | 143P | 2000-12-27 |  | 1047.36 | 643.6 | -0.1 | 152.8774 | 269.4585 | 6.0336
7/12 | 111P | 2000-12-27 |  | 839.08 | 643.6 | -0.1 | 142.9474 | 230.9804 | 5.172
7/11 | 112P | 2000-12-27 |  | 830.3 | 643.6 | -0.1 | 140.726 | 232.2179 | 5.1997
7/9 | 123P | 2000-12-27 |  | 888.56 | 643.6 | -0.1 | 147.3611 | 237.2033 | 5.3113
7/7 | 108P | 2000-12-27 |  | 939.23 | 643.6 | -0.1 | 139.7139 | 264.6827 | 5.9266
7/5 | 120P | 2000-12-27 |  | 996.31 | 643.6 | -0.1 | 145.3855 | 269.6931 | 6.0388
7/3 | 141P | 2000-12-27 |  | 996.55 | 643.6 | -0.1 | 142.9054 | 274.5062 | 6.1466
7/1 | 109P | 2000-12-27 |  | 980.87 | 643.6 | -0.1 | 140.5172 | 274.8373 | 6.154
LMG 02 03 (04 2002)
Test | [9 empty] | N. Bottle #
4/13 2002 one depth | 1247 | 2002-04-14 |  | 1256.31 | 644.9 | -0.1 | 144.8102 | 340.8852 | 7.6329 | 1 | *****
| 1254 | 2002-04-14 |  | 1157.35 | 644.9 | -0.1 | 145.3896 | 312.725 | 7.0024 | 2 | *****
| 1059 | 2002-04-14 |  | 1248.71 | 644.9 | -0.1 | 145.1145 | 338.0996 | 7.5705 | 3
| 1244 | 2002-04-14 |  | 1269.87 | 644.9 | -0.1 | 147.5013 | 338.1971 | 7.5727 | 4
| 1255 | 2002-04-14 |  | 1227.86 | 644.9 | -0.1 | 143.1182 | 337.1487 | 7.5492 | 5
| 1275 | 2002-04-14 |  | 1231.28 | 644.9 | -0.1 | 143.1527 | 338.0066 | 7.5684 | 6
| 1208 | 2002-04-14 |  | 1255.27 | 644.9 | -0.1 | 145.9126 | 337.9935 | 7.5681 | 7
| 1211 | 2002-04-14 |  | 1169.34 | 644.9 | -0.1 | 136.1809 | 337.6547 | 7.5606 | 8
| 1262 | 2002-04-14 |  | 1236.03 | 644.9 | -0.1 | 144.4554 | 336.2096 | 7.5282 | 9
| 1216 | 2002-04-14 |  | 1252.48 | 644.9 | -0.1 | 146.3039 | 336.3267 | 7.5308 | 10
| 1263 | 2002-04-14 |  | 1238.43 | 644.9 | -0.1 | 145.0107 | 335.5554 | 7.5136 | 11
| 1273 | 2002-04-14 |  | 1229.2 | 644.9 | -0.1 | 143.7998 | 335.8948 | 7.5212 | 12
| 1256 | 2002-04-14 |  | 1239.11 | 644.9 | -0.1 | 144.8434 | 336.1331 | 7.5265 | 13
| 1198 | 2002-04-14 |  | 1242.51 | 644.9 | -0.1 | 145.4402 | 335.6545 | 7.5158 | 14
| 1239 | 2002-04-14 |  | 1248.4 | 644.9 | -0.1 | 145.9792 | 335.9855 | 7.5232 | 15
| 1040 | 2002-04-14 |  | 1246.58 | 644.9 | -0.1 | 146.0074 | 335.4292 | 7.5107 | 16
| 1271 | 2002-04-14 |  | 1231.44 | 644.9 | -0.1 | 144.1623 | 335.6498 | 7.5157 | 17
| 1249 | 2002-04-14 |  | 1269.22 | 644.9 | -0.1 | 148.5137 | 335.688 | 7.5165 | 18
| 1253 | 2002-04-14 |  | 1244.01 | 644.9 | -0.1 | 145.4709 | 335.9884 | 7.5233 | 19
| 1289 | 2002-04-14 |  | 1232.04 | 644.9 | -0.1 | 144.2637 | 335.5742 | 7.514 | 20
| 1270 | 2002-04-14 |  | 1235.3 | 644.9 | -0.1 | 144.4504 | 336.0225 | 7.524 | 21
| 1290 | 2002-04-14 |  | 1288.57 | 644.9 | -0.1 | 143.4339 | 353.054 | 7.9054 | 22 | *******
| 1217 | 2002-04-14 |  | 1158.14 | 644.9 | -0.1 | 135.5094 | 336.0975 | 7.5257 | 23
| 1130 | 2002-04-14 |  | 1244.18 | 644.9 | -0.1 | 145.1189 | 336.8608 | 7.5428 | 24
LMG 02 03 (04 2002)
Event 19 |  | 2002-04-15 | [7 empty] | Niskin # | Depth (m)
| 1254 | 2002-04-15 |  | 675.55 | 647.3 | -0.1 | 145.3896 | 181.6516 | 4.0674 | 1 | 465
| 1247 | 2002-04-15 |  | 697.49 | 647.3 | -0.1 | 144.8102 | 188.3285 | 4.2169 | 6 | 153
| 1244 | 2002-04-15 |  | 814.09 | 647.3 | -0.1 | 147.5013 | 215.8289 | 4.8327 | 7 | 100.4
| 1059 | 2002-04-15 |  | 980.71 | 647.3 | -0.1 | 145.1145 | 264.442 | 5.9212 | 9 | 46
| 1255 | 2002-04-15 |  | 1222.81 | 647.3 | -0.1 | 143.1182 | 334.5133 | 7.4902 | 11 | 2
Event 31
| 1262 | 2002-04-15 |  | 671.84 | 647.3 | -0.1 | 144.4554 | 181.8359 | 4.0716 | 1 | 239.7
| 1275 | 2002-04-15 |  | 809.36 | 647.3 | -0.1 | 143.1527 | 221.1825 | 4.9526 | 5 | 92
| 1216 | 2002-04-15 |  | 1255.83 | 647.3 | -0.1 | 146.3039 | 335.9755 | 7.523 | 9 | 24
| 1208 | 2002-04-15 |  | 1238.19 | 647.3 | -0.1 | 145.9126 | 332.1499 | 7.4373 | 11 | 2
Event 39 Cast 12 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m) | ctd
| 1130 | 2002-04-16 | 2002-04-17 | 697.72 | 645 | -0.1 | 145.1189 | 188.6565 | 4.2243 | 3 | 785 | 4.388
| 1270 | 2002-04-16 |  | 685.8 | 645 | -0.1 | 144.4504 | 186.295 | 4.1714 | 5 | 524 | 4.305
| 1208 | 2002-04-16 |  | 687.65 | 645 | -0.1 | 145.9126 | 184.901 | 4.1402 | 8 | 341 | 4.244
| 1216 | 2002-04-16 |  | 617.62 | 645 | -0.1 | 146.3039 | 165.5696 | 3.7073 | 9 | 185 | 3.8203
| 1262 | 2002-04-16 |  | 641.7 | 645 | -0.1 | 144.4554 | 174.2768 | 3.9023 | 11 | 146 | 4.012
| 1275 | 2002-04-16 |  | 643.23 | 645 | -0.1 | 143.1527 | 176.3058 | 3.9477 | 13 | 129 | 4.052
| 1059 | 2002-04-16 |  | 840.1 | 645 | -0.1 | 145.1145 | 227.2645 | 5.0888 | 15 | 83 | 5.134
| 1244 | 2002-04-16 |  | 1265.12 | 645 | -0.1 | 147.5013 | 336.8779 | 7.5432 | 20 | 2 | 7.559
Event 51 cast 14
| 1255 | 2002-04-17 |  | 687.46 | 645 | -0.1 | 143.1182 | 188.5102 | 4.221 | 1 | 752
| 1211 | 2002-04-17 |  | 637.34 | 645 | -0.1 | 136.1809 | 183.7631 | 4.1147 | 5 | 390
| 1198 | 2002-04-17 |  | 662.22 | 645 | -0.1 | 145.4402 | 178.631 | 3.9998 | 7 | 305
| 1040 | 2002-04-17 |  | 613.71 | 645 | -0.1 | 146.0074 | 164.8569 | 3.6914 | 9 | 206
| 1273 | 2002-04-17 |  | 849.23 | 645 | -0.1 | 143.7998 | 231.8699 | 5.1919 | 11 | 71
| 1256 | 2002-04-17 |  | 1247.67 | 645 | -0.1 | 144.8434 | 338.4061 | 7.5774 | 13 | 2
Event 70 Cast 15 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1255 | 2002-04-19 |  | 615.29 | 644.44 | -0.1 | 143.1182 | 168.8136 | 3.78 | 1 | 354
| 1208 | 2002-04-19 |  | 625.67 | 644.44 | -0.1 | 145.9126 | 168.3367 | 3.7693 | 5 | 193
| 1239 | 2002-04-19 |  | 708.26 | 644.44 | -0.1 | 145.9792 | 190.5356 | 4.2664 | 7 | 121
| 1247 | 2002-04-19 |  | 1143.94 | 644.44 | -0.1 | 144.8102 | 310.5716 | 6.9541 | 13 | 31
| 1262 | 2002-04-19 |  | 1207.57 | 644.44 | -0.1 | 144.4554 | 328.6914 | 7.3599 | 15 | 2
Event 86 Cast 16 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1211 | 2002-04-19 |  | 615.68 | 644.44 | -0.1 | 136.1809 | 177.6543 | 3.9779 | 1 | 167
| 1290 | 2002-04-19 |  | 817.16 | 644.44 | -0.1 | 143.4339 | 223.8664 | 5.0127 | 5 | 73
| 1216 | 2002-04-19 |  | 1168.33 | 644.44 | -0.1 | 146.3039 | 313.9207 | 7.0291 | 7 | 53
| 1275 | 2002-04-19 |  | 1215.74 | 644.44 | -0.1 | 143.1527 | 333.9727 | 7.4781 | 9 | 2
Event 92 Cast 18 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1059 | 2002-04-19 |  | 610.4 | 644.44 | -0.1 | 145.1145 | 165.1319 | 3.6975 | 1 | 240
| 1198 | 2002-04-19 |  | 631.13 | 644.44 | -0.1 | 145.4402 | 170.3694 | 3.8148 | 3 | 179
| 1273 | 2002-04-19 |  | 644.94 | 644.44 | -0.1 | 143.7998 | 176.1225 | 3.9436 | 5 | 144
| 1244 | 2002-04-19 |  | 1042.87 | 644.44 | -0.1 | 147.5013 | 277.8514 | 6.2215 | 9 | 70
| 1271 | 2002-04-19 |  | 1228.11 | 644.44 | -0.1 | 144.1623 | 334.9801 | 7.5007 | 11 | 2
Event 100 Cast 19 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1208 | 2002-04-23 |  | 612.21 | 645.8533 | -0.1 | 145.9126 | 164.343 | 3.6799 | 3 | 701
| 1255 | 2002-04-23 |  | 625.95 | 645.8533 | -0.1 | 143.1182 | 171.3702 | 3.8372 | 7 | 320
| 1239 | 2002-04-23 |  | 622.68 | 645.8533 | -0.1 | 145.9792 | 167.0849 | 3.7413 | 9 | 239
| 1247 | 2002-04-23 |  | 620.05 | 645.8533 | -0.1 | 144.8102 | 167.7389 | 3.7559 | 11 | 201
| 1216 | 2002-04-23 |  | 757.67 | 645.8533 | -0.1 | 146.3039 | 202.9579 | 4.5445 | 13 | 115
| 1211 | 2002-04-23 |  | 1114.09 | 645.8533 | -0.1 | 136.1809 | 321.1999 | 7.1921 | 15 | 70
| 1290 | 2002-04-23 |  | 1281.15 | 645.8533 | -0.1 | 143.4339 | 350.4993 | 7.8482 | 17 | 2
Event 112 Cast 20 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1208 | 2002-04-24 |  | 808.56 | 645.8533 | -0.1 | 145.9126 | 217.2124 | 4.8637 | 1 | 2942
| 1239 | 2002-04-24 |  | 786.32 | 645.8533 | -0.1 | 145.9792 | 211.1263 | 4.7274 | 5 | 1984
| 1256 | 2002-04-24 |  | 736.79 | 645.8533 | -0.1 | 144.8434 | 199.3687 | 4.4641 | 11 | 988
| 1247 | 2002-04-24 |  | 685.56 | 645.8533 | -0.1 | 144.8102 | 185.5144 | 4.1539 | 15 | 541
| 1216 | 2002-04-24 |  | 673.6 | 645.8533 | -0.1 | 146.3039 | 180.3825 | 4.039 | 16 | 361
| 1211 | 2002-04-24 |  | 620.59 | 645.8533 | -0.1 | 136.1809 | 178.6823 | 4.0009 | 17 | 246
| 1262 | 2002-04-24 |  | 1295.43 | 645.8533 | -0.1 | 144.4554 | 351.8703 | 7.8789 | 21 | 2
Event 122 Cast 21 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1217 | 2002-04-24 |  | 749.24 | 645.8533 | -0.1 | 135.5094 | 216.9207 | 4.8572 | 1 | 3136
| 1198 | 2002-04-24 |  | 736.72 | 645.8533 | -0.1 | 145.4402 | 198.5203 | 4.4451 | 7 | 1600
| 1059 | 2002-04-24 |  | 683.61 | 645.8533 | -0.1 | 145.1145 | 184.5919 | 4.1333 | 8 | 500
| 1275 | 2002-04-24 |  | 645.78 | 645.8533 | -0.1 | 143.1527 | 176.7722 | 3.9582 | 10 | 251
| 1244 | 2002-04-24 |  | 1337.45 | 645.8533 | -0.1 | 147.5013 | 355.6951 | 7.9645 | 14 | 2
Event 132 Cast 22 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1208 | 2002-04-29 |  | 710.21 | 645 | -0.1 | 145.9126 | 190.9836 | 4.2764 | 1 | 497
| 1216 | 2002-04-29 |  | 675.04 | 645 | -0.1 | 146.3039 | 181.009 | 4.053 | 3 | 424
| 1059 | 2002-04-29 |  | 663.78 | 645 | -0.1 | 145.1145 | 179.4605 | 4.0184 | 5 | 300
| 1262 | 2002-04-29 |  | 652.28 | 645 | -0.1 | 144.4554 | 177.1585 | 3.9668 | 7 | 193
| 1198 | 2002-04-29 |  | 1057.57 | 645 | -0.1 | 145.4402 | 285.5752 | 6.3944 | 11 | 87
| 1275 | 2002-04-29 |  | 1235.85 | 645 | -0.1 | 143.1527 | 339.2103 | 7.5954 | 13 | 60
| 1244 | 2002-04-29 |  | 1335.96 | 645 | -0.1 | 147.5013 | 355.769 | 7.9662 | 15 | 2
Event 144 Cast 25 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1256 | 2002-05-02 |  | 676.48 | 646.8 | -0.1 | 144.8434 | 182.7395 | 4.0918 | 1 | 738
| 1275 | 2002-05-02 |  | 667.27 | 646.8 | -0.1 | 143.1527 | 182.4036 | 4.0843 | 3 | 550
| 1198 | 2002-05-02 |  | 670.57 | 646.8 | -0.1 | 145.4402 | 180.3849 | 4.0391 | 5 | 294
| 1040 | 2002-05-02 |  | 815.76 | 646.8 | -0.1 | 146.0074 | 218.6855 | 4.8967 | 7 | 126
| 1208 | 2002-05-02 |  | 1304.54 | 646.8 | -0.1 | 145.9126 | 350.2463 | 7.8425 | 9 | 42
| 1130 | 2002-05-02 |  | 1283.17 | 646.8 | -0.1 | 145.1189 | 346.4111 | 7.7566 | 11 | 2
Event 154 Cast 26 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1262 | 2002-05-02 |  | 684.67 | 646.8 | -0.1 | 144.4554 | 185.4617 | 4.1527 | 1 | 991
| 1211 | 2002-05-02 |  | 638.12 | 646.8 | -0.1 | 136.1809 | 183.4751 | 4.1083 | 3 | 667
| 1247 | 2002-05-02 |  | 669.4 | 646.8 | -0.1 | 144.8102 | 180.8637 | 4.0498 | 5 | 310
| 1217 | 2002-05-02 |  | 637.59 | 646.8 | -0.1 | 135.5094 | 184.2443 | 4.1255 | 7 | 192
| 1239 | 2002-05-02 |  | 953.09 | 646.8 | -0.1 | 145.9792 | 255.6348 | 5.724 | 9 | 95
| 1290 | 2002-05-02 |  | 1265.18 | 646.8 | -0.1 | 143.4339 | 345.6165 | 7.7388 | 11 | 56
| 1273 | 2002-05-02 |  | 1262.74 | 646.8 | -0.1 | 143.7998 | 344.0589 | 7.704 | 13 | 20
| 1271 | 2002-05-02 |  | 1289.8 | 646.8 | -0.1 | 144.1623 | 350.5467 | 7.8492 | 15 | 2
Event 166 Cast 27 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1256 | 2002-05-05 |  | 675.85 | 645.5 | -0.1 | 144.8434 | 182.9375 | 4.0962 | 1 | 790
| 1211 | 2002-05-05 |  | 631.45 | 645.5 | -0.1 | 136.1809 | 181.9184 | 4.0734 | 4 | 500
| 1247 | 2002-05-05 |  | 664.53 | 645.5 | -0.1 | 144.8102 | 179.9068 | 4.0284 | 7 | 230
| 1275 | 2002-05-05 |  | 675.06 | 645.5 | -0.1 | 143.1527 | 184.9117 | 4.1404 | 9 | 181
| 1262 | 2002-05-05 |  | 802.75 | 645.5 | -0.1 | 144.4554 | 217.9734 | 4.8807 | 11 | 137
| 1198 | 2002-05-05 |  | 1158.14 | 645.5 | -0.1 | 145.4402 | 312.5372 | 6.9981 | 16 | 47
| 1130 | 2002-05-05 |  | 1274.06 | 645.5 | -0.1 | 145.1189 | 344.6418 | 7.717 | 19 | 2
Event 174 Cast 28 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1040 | 2002-05-05 |  | 680.39 | 645.5 | -0.1 | 146.0074 | 182.6812 | 4.0905 | 1 | 639
| 1208 | 2002-05-05 |  | 672.79 | 645.5 | -0.1 | 145.9126 | 180.754 | 4.0473 | 5 | 399
| 1217 | 2002-05-05 |  | 634.68 | 645.5 | -0.1 | 135.5094 | 183.7714 | 4.1149 | 7 | 180
| 1273 | 2002-05-05 |  | 986.15 | 645.5 | -0.1 | 143.7998 | 269.1269 | 6.0261 | 9 | 81
| 1290 | 2002-05-05 |  | 1277.41 | 645.5 | -0.1 | 143.4339 | 349.6661 | 7.8295 | 11 | 39
| 1239 | 2002-05-05 |  | 1324.2 | 645.5 | -0.1 | 145.9792 | 356.0844 | 7.9732 | 13 | 2
Event 178 Cast 29 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1211 | 2002-05-08 |  | 635.66 | 647.1 | -0.1 | 136.1809 | 182.6807 | 4.0905 | 1 | 408
| 1247 | 2002-05-08 |  | 684.17 | 647.1 | -0.1 | 144.8102 | 184.7796 | 4.1375 | 5 | 274
| 1270 | 2002-05-08 |  | 748.82 | 647.1 | -0.1 | 144.4504 | 202.7988 | 4.5409 | 7 | 143
| 1244 | 2002-05-08 |  | 981.69 | 647.1 | -0.1 | 147.5013 | 260.4452 | 5.8317 | 11 | 58
| 1275 | 2002-05-08 |  | 1138.86 | 647.1 | -0.1 | 143.1527 | 311.533 | 6.9757 | 13 | 36
| 1040 | 2002-05-08 |  | 1246.87 | 647.1 | -0.1 | 146.0074 | 334.3651 | 7.4869 | 15 | 0.3
Event 190 Cast 30 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1130 | 2002-05-08 |  | 756.47 | 647.1 | -0.1 | 145.1189 | 203.9188 | 4.566 | 2 | 260
| 1262 | 2002-05-08 |  | 755.28 | 647.1 | -0.1 | 144.4554 | 204.5455 | 4.5801 | 4 | 201
| 1198 | 2002-05-08 |  | 891.57 | 647.1 | -0.1 | 145.4402 | 239.8887 | 5.3714 | 8 | 114
| 1273 | 2002-05-08 |  | 1024.36 | 647.1 | -0.1 | 143.7998 | 278.8818 | 6.2446 | 10 | 91
| 1290 | 2002-05-08 |  | 1220.68 | 647.1 | -0.1 | 143.4339 | 333.2874 | 7.4628 | 14 | 55
| 1217 | 2002-05-08 |  | 1227.73 | 647.1 | -0.1 | 135.5094 | 355.1121 | 7.9515 | 18 | 0
Event 193 Cast 31 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1217 | 2002-05-10 |  | 582.1 | 648.3 | -0.1 | 135.5094 | 167.772 | 3.7567 | 1 | 446
| 1262 | 2002-05-10 |  | 613.42 | 648.3 | -0.1 | 144.4554 | 165.7235 | 3.7108 | 5 | 233
| 1290 | 2002-05-10 |  | 607.57 | 648.3 | -0.1 | 143.4339 | 165.3237 | 3.7018 | 6 | 148
| 1270 | 2002-05-10 |  | 780.64 | 648.3 | -0.1 | 144.4504 | 211.0457 | 4.7256 | 8 | 80
| 1211 | 2002-05-10 |  | 1061.89 | 648.3 | -0.1 | 136.1809 | 304.9679 | 6.8287 | 11 | 38
| 1244 | 2002-05-10 |  | 1292.48 | 648.3 | -0.1 | 147.5013 | 342.4197 | 7.6673 | 15 | 0.5
Event 201 Cast 32 | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | Niskin # | Depth (m)
| 1198 | 2002-05-10 |  | 596.48 | 648.3 | -0.1 | 145.4402 | 160.0267 | 3.5832 | 1 | 224
| 1040 | 2002-05-10 |  | 614.07 | 648.3 | -0.1 | 146.0074 | 164.1117 | 3.6747 | 6 | 150
| 1273 | 2002-05-10 |  | 716.08 | 648.3 | -0.1 | 143.7998 | 194.4381 | 4.3537 | 8 | 100
| 1130 | 2002-05-10 |  | 1047.2 | 648.3 | -0.1 | 145.1189 | 281.9598 | 6.3135 | 12 | 65
| 1059 | 2002-05-10 |  | 1229.92 | 648.3 | -0.1 | 145.1145 | 331.2555 | 7.4173 | 15 | 39
| 1275 | 2002-05-10 |  | 1262.17 | 648.3 | -0.1 | 143.1527 | 344.6794 | 7.7179 | 17 | 0.5

## BotVols
|  |  |  | °C | rw |  | fbuoy
|  |  | DIW | 24 | 0.9973 | Buoyancy correction | 1.001
Bottle No. | Empty | Full | Vol (24C) | Vol(20C)
100P | 84.452 | 227.2445 | 143.3295 | 143.3238
101P | 84.952 | 225.3775 | 140.9536 | 140.948
102P | 85.402 | 224.5637 | 139.6851 | 139.6795
103P | 84.2388 | 226.2134 | 142.5085 | 142.5028
104P | 84.2286 | 228.7372 | 145.0521 | 145.0463
105P | 84.1618 | 223.4192 | 139.7811 | 139.7755
106P | 85.9912 | 223.0814 | 137.6058 | 137.6003
107P | 85.0865 | 225.4279 | 140.8692 | 140.8636
108P | 84.7133 | 223.9093 | 139.7195 | 139.7139
109P | 85.5466 | 225.5429 | 140.5228 | 140.5172
110P | 85.4673 | 224.9417 | 139.9989 | 139.9933
111P | 84.5597 | 226.9772 | 142.9531 | 142.9474
112P | 85.2297 | 225.4341 | 140.7317 | 140.726
113P | 83.8293 | 217.5013 | 134.1747 | 134.1693
114P | 83.8244 | 223.8 | 140.502 | 140.4964
115P | 84.54 | 223.0199 | 139.0007 | 138.9951
116P | 86.0119 | 226.9849 | 141.5032 | 141.4975
117P | 85.3855 | 225.3237 | 140.4645 | 140.4589
118P | 85.4802 | 227.9381 | 142.9936 | 142.9879
119P | 85.6576 | 225.9532 | 140.8232 | 140.8176
120P | 83.2334 | 228.08 | 145.3913 | 145.3855
121P | 85.7383 | 226.5056 | 141.2967 | 141.291 | 0.004
122P | 84.5367 | 227.2735 | 143.2736 | 143.2679
123P | 85.4492 | 232.2641 | 147.367 | 147.3611
124P | 84.936 | 234.6342 | 150.2612 | 150.2552
125P | 85.4683 | 227.1948 | 142.2595 | 142.2538
126P | 85.4495 | 223.4206 | 138.49 | 138.4844
127P | 84.6305 | 216.8697 | 132.7365 | 132.7312
128P | 85.1837 | 233.1586 | 148.5314 | 148.5255
129P | 85.6474 | 227.1187 | 142.0033 | 141.9977
130P | 84.5383 | 221.6705 | 137.6479 | 137.6424
131P | 85.008 | 228.2599 | 143.7906 | 143.7849
132P | 86.3757 | 228.0737 | 142.2309 | 142.2252
133P | 85.1368 | 226.8489 | 142.245 | 142.2394
134P | 84.6774 | 223.5234 | 139.3682 | 139.3626
135P | 84.3653 | 220.5615 | 136.7084 | 136.7029
136P | 86.5055 | 225.7143 | 139.7323 | 139.7267
137P | 86.2536 | 228.134 | 142.414 | 142.4083
138P | 84.4294 | 223.1251 | 139.2173 | 139.2117
139P | 87.2229 | 232.7799 | 146.1044 | 146.0986
140P | 85.1274 | 223.0346 | 138.4258 | 138.4203
141P | 88.5494 | 230.9251 | 142.9111 | 142.9054
142P | 86.627 | 226.6021 | 140.5015 | 140.4959
143P | 86.2936 | 238.6043 | 152.8835 | 152.8774
144P | 85.5717 | 237.2192 | 152.2178 | 152.2117
145P | 85.8435 | 228.293 | 142.9852 | 142.9795
146P | 86.0746 | 226.7127 | 141.167 | 141.1614
147P | 86.4491 | 229.2448 | 143.3327 | 143.327
148P | 85.9547 | 224.274 | 138.8395 | 138.8339
149P | 83.113 | 223.9686 | 141.3853 | 141.3797
150P | 84.3232 | 225.3355 | 141.5426 | 141.5369
151P | 86.0976 | 232.9885 | 147.4433 | 147.4374
152P | 84.8629 | 221.8124 | 137.4645 | 137.459
153P | 85.3353 | 225.9817 | 141.1753 | 141.1697
154P | 84.4189 | 226.8484 | 142.9651 | 142.9594
155P | 86.599 | 226.104 | 140.0296 | 140.024
156P | 85.6404 | 226.833 | 141.7236 | 141.7179
157P | 83.7576 | 229.3419 | 146.1318 | 146.126
1000 | 84.277 | 228.293 |  | 144.5291
1001 | 84.13 | 229.437 |  | 145.8247
1002 | 84.614 | 228.347 |  | 144.2451
1003 | 85.123 | 228.722 |  | 144.114
1004 | 89.938 | 194.782 |  | 105.3146
1005 | 94.059 | 198.262 |  | 104.6707
1006 | 91.882 | 193.847 |  | 102.4227
1007 | 91.089 | 194.061 |  | 103.4342
1008 | 92.302 | 195.803 |  | 103.9655
1009 | 85.145 | 183.216 |  | 98.5112
1010 | 93.137 | 194.197 |  | 101.5136
1011 | 93.299 | 191.651 |  | 98.7934
1012 | 82.688 | 192.604 |  | 110.4093
1013 | 95.66 | 198.741 |  | 103.5464
1014 | 98.139 | 199.772 |  | 102.0919
1015 | 99.019 | 197.994 |  | 99.4219
1016 | 86.555 | 198.69 |  | 112.6383
1017 | 89.087 | 234.758 |  | 146.2459
1018 | 87.586 | 235.246 |  | 148.2428
1019 | 88.502 | 233.259 |  | 145.3283
1020 | 89.038 | 235.138 |  | 146.6766
1021 | 85.582 | 230.572 |  | 145.5623
1022 | 87.788 | 232.215 |  | 144.997
1023 | 86.859 | 233.356 |  | 147.0788
1024 | 86.839 | 232.379 |  | 146.118
1025 | 87.004 | 234.726 |  | 148.3087
1026 | 87.608 | 235.58 |  | 148.4722
1027 | 86.125 | 232.493 |  | 146.8628
1028 | 84.505 | 231.245 |  | 147.2361
1029 | 85.058 | 228.951 |  | 144.3794
1030 | 86.2 | 230.428 |  | 144.7156
1031 | 84.329 | 229.809 |  | 145.9718
1032 | 85.027 | 230.768 |  | 146.2337
1033 | 86.202 | 231.825 |  | 146.1185
#1034 | 84.91 | 230.105 |  | 145.6841
1034 | 89.372 | 234.475 |  | 145.6476
1035 | 85.761 | 230.628 |  | 145.36
1036 | 83.604 | 227.443 |  | 144.3253
1037 | 87.291 | 234.875 |  | 148.0929
1038 | 86.732 | 232.596 |  | 146.4043
1039 | 88.109 | 232.694 |  | 145.1171
1040 | 83.472 | 228.944 |  | 146.0074
1041 | 87.794 | 231.367 |  | 144.1014
1042 | 87.035 | 230.64 |  | 144.1369
1043 | 87.703 | 234.057 |  | 146.8961
1044 | 86.013 | 229.914 |  | 144.4306
1045 | 85.377 | 230.219 |  | 145.3785
1046 | 87.785 | 231.979 |  | 144.735
1047 | 88.27 | 232.392 |  | 144.6627
1048 | 87.702 | 232.405 |  | 145.2459
1049 | 87.621 | 233.643 |  | 146.5698
1050 | 86.233 | 228.256 |  | 142.5192
1051 | 86.167 | 230.442 |  | 144.7758
1052 | 83.019 | 225.422 |  | 142.8973
1053 | 87.217 | 233.212 |  | 146.5051
1054 | 87.157 | 232.454 |  | 145.8046
1055 | 86.312 | 230.842 |  | 145.035
1056 | 84.904 | 229.311 |  | 144.9115
1057 | 84.669 | 229.572 |  | 145.4093
1058 | 87.662 | 233.417 |  | 146.2676
1059 | 84.976 | 229.582 |  | 145.1145
1060 | 85.926 | 231.186 |  | 145.7708
1061 | 89.279 | 234.846 |  | 146.0789
1062 | 87.373 | 230.767 |  | 143.8983
1063 | 86.023 | 230.977 |  | 145.4638
1064 | 88.929 | 233.467 |  | 145.0496
1065 | 83.511 | 230.508 |  | 147.5139
1066 | 84.131 | 227.812 |  | 144.1896
1067 | 86.063 | 233.812 |  | 148.272
1068 | 89.873 | 232.811 |  | 143.444
1069 | 84.158 | 228.339 |  | 144.6914
1070 | 88.29 | 237.239 |  | 149.4797
1071 | 83.376 | 231.949 |  | 149.1024
1072 | 88.737 | 233.236 |  | 145.0005
1073 | 88.303 | 237.325 |  | 149.5392
1074 | 87.036 | 232.35 |  | 145.8385
1075 | 91.897 | 239.149 |  | 147.7082
1076 | 87.229 | 231.775 |  | 145.0711
1077 | 83.181 | 229.799 |  | 147.1506
1078 | 85.42 | 231.173 |  | 146.2825
1079 | 85.157 | 230.562 |  | 145.9333
1080 | 86.484 | 234.775 |  | 148.8297
1081 | 84.192 | 229.184 |  | 145.5188
1082 | 84.993 | 232.959 |  | 148.507
1083 | 84.798 | 229.12 |  | 144.8497
1084 | 83.097 | 230.325 |  | 147.7698
1085 | 87.301 | 232.658 |  | 145.8885
1086 | 85.841 | 230.04 |  | 144.7297
1087 | 86.824 | 231.285 |  | 144.9926
1088 | 84.675 | 232.025 |  | 147.8923
1089 | 88.453 | 230.534 |  | 142.6039
1090 | 84.248 | 228.496 |  | 144.7788
1091 | 88.059 | 233.114 |  | 145.5888
1092 | 86.725 | 231.734 |  | 145.5427
1093 | 85.424 | 232.392 |  | 147.5089
1094 | 86.193 | 228.452 |  | 142.7825
1095 | 86.597 | 232.15 |  | 146.0886
1096 | 87.196 | 230.772 |  | 144.1044
1097 | 86.963 | 230.636 |  | 144.2018
1098 | 83.847 | 227.413 |  | 144.1012
1099 | 86.656 | 232.567 |  | 146.4619
1100 | 84.998 | 231.395 |  | 146.9533
1101 | 83.984 | 231.28 |  | 147.8521
1102 | 88.246 | 231.498 |  | 143.7929
1103 | 84.423 | 227.506 |  | 143.6233
1104 | 85.907 | 230.782 |  | 145.422
1105 | 86.537 | 232.311 |  | 146.3244
1106 | 86.516 | 232.85 |  | 146.8353
1107 | 85.364 | 232.373 |  | 147.5126
1108 | 88.16 | 233.408 |  | 145.7456
1109 | 86.422 | 233.666 |  | 147.7451
1110 | 85.844 | 233.977 |  | 148.6371
1111 | 83.935 | 227.299 |  | 143.8519
1112 | 86.805 | 230.226 |  | 143.9091
1113 | 86.631 | 231.889 |  | 145.7523
1114 | 85.075 | 229.354 |  | 144.77
1115 | 86.341 | 233.033 |  | 147.1912
1116 | 86.929 | 234.67 |  | 148.2404
1117 | 87.203 | 231.977 |  | 145.2667
1118 | 84.72 | 229.378 |  | 145.1503
1119 | 88.078 | 231.687 |  | 144.0977
1120 | 85.96 | 232.535 |  | 147.0738
1121 | 84.74 | 229.837 |  | 145.5908
1122 | 86.653 | 235.43 |  | 149.2015
1123 | 86.18 | 233.115 |  | 147.3542
1124 | 81.568 | 226.815 |  | 145.6614
1125 | 84.679 | 228.47 |  | 144.1984
1126 | 87.166 | 233.057 |  | 146.3044
1127 | 88.478 | 233.604 |  | 145.5372
1128 | 83.432 | 230.66 |  | 147.651
1129 | 88.372 | 232.366 |  | 144.4048
1130 | 86.945 | 231.651 |  | 145.1189
1131 | 87.275 | 235.205 |  | 148.3521
1132 | 86.645 | 233.552 |  | 147.3262
1133 | 84.644 | 230.721 |  | 146.4967
1134 | 84.068 | 231.095 |  | 147.4672
1135 | 82.657 | 229.679 |  | 147.4592
1136 | 86.086 | 232.281 |  | 146.6298
1137 | 85.701 | 232.227 |  | 146.9617
1138 | 87.071 | 233.496 |  | 146.8604
1139 | 86.091 | 234.174 |  | 148.5264
1140 | 86.801 | 229.459 |  | 143.0851
1141 | 87.934 | 233.583 |  | 146.0851
1142 | 85.927 | 233.619 |  | 148.0877
1143 | 86.172 | 230.72 |  | 144.9326
1144 | 88.236 | 236.917 |  | 149.0766
1145 | 93.327 | 241.025 |  | 148.1842
1146 | 86.096 | 232.528 |  | 146.8216
1147 | 85.56 | 230.098 |  | 144.9253
1148 | 85.31 | 230.172 |  | 145.2529
1149 | 84.317 | 232.271 |  | 148.3532
1150 | 85.188 | 234.769 |  | 149.9903
1151 | 87.787 | 231.749 |  | 144.3559
1152 | 86.405 | 229.463 |  | 143.4495
1153 | 89.649 | 233.081 |  | 143.8272
1154 | 87.237 | 228.869 |  | 142.0223
1155 | 85.448 | 233.011 |  | 147.9696
1156 | 87.339 | 240.745 |  | 153.8287
1157 | 87.484 | 233.71 |  | 146.629
1158 | 87.999 | 234.295 |  | 146.7221
1159 | 86.241 | 235.148 |  | 149.3408
1160 | 86.272 | 231.692 |  | 145.8436
1161 | 86.484 | 231.693 |  | 145.632
1162 | 83.385 | 218.212 |  | 135.2197
1163 | 84.317 | 233.837 |  | 149.9556
1164 | 85.629 | 228.462 |  | 143.2491
1165 | 87.26 | 232.284 |  | 145.4464
1166 | 86.313 | 232.097 |  | 146.2087
1167 | 81.095 | 222.941 |  | 142.2423
1168 | 86.047 | 234.12 |  | 148.4839
1169 | 86.889 | 230.259 |  | 143.7678
1170 | 85.284 | 231.608 |  | 146.7301
1171 | 87.372 | 235.02 |  | 148.0577
1172 | 86.607 | 234.884 |  | 148.6913
1173 | 83.305 | 221.863 |  | 138.9452
1174 | 84.734 | 233.4 |  | 149.0814
1175 | 84.903 | 230.213 |  | 145.716
1176 | 85.824 | 230.345 |  | 144.9248
1177 | 84.385 | 234.013 |  | 150.0461
1178 | 88.577 | 230.578 |  | 142.4118
1179 | 85.405 | 231.519 |  | 146.5367
1180 | 86.991 | 231.155 |  | 144.5811
1181 | 86.491 | 234.89 |  | 148.8283
1182 | 83.561 | 228.483 |  | 145.3412
1183 | 84.923 | 228.281 |  | 143.7756
1184 | 87.222 | 235.666 |  | 148.8764
1185 | 86.562 | 221.382 |  | 135.2154
1186 | 84.993 | 232.46 |  | 147.8995
1187 | 84.207 | 227.862 |  | 144.0793
1188 | 85.538 | 229.657 |  | 144.5446
1189 | 86.281 | 234.569 |  | 148.7259
1190 | 87.765 | 220.943 |  | 133.5713
1191 | 89.428 | 233.84 |  | 144.8385
1192 | 84.563 | 222.154 |  | 137.9974
1193 | 83.84 | 227.7 |  | 144.2849
1194 | 85.522 | 231.316 |  | 146.2276
1195 | 114.924 | 249.282 |  | 134.7576
1196 | 85.12 | 230.105 |  | 145.4132
1197 | 82.821 | 227.19 |  | 144.7983
1198 | 85.031 | 230.04 |  | 145.4402
1199 | 84.47 | 230.773 |  | 146.7381
1200 | 84.446 | 230.242 |  | 146.2386
1201 | 82.453 | 228.092 |  | 146.078
1202 | 81.933 | 228.24 |  | 146.7481
1203 | 85.105 | 230.57 |  | 145.9035
1204 | 82.164 | 228.545 |  | 146.8223
1205 | 114.785 | 247.703 |  | 133.3187
1206 | 84.706 | 233.533 |  | 149.2819
1207 | 86.047 | 231.358 |  | 145.7551
1208 | 84.664 | 230.132 |  | 145.9126
1209 | 85.645 | 231.511 |  | 146.3118
1210 | 86.553 | 231.207 |  | 145.0961
1211 | 83.883 | 219.649 |  | 136.1809
1212 | 88.051 | 234.573 |  | 146.9759
1213 | 85.741 | 230.955 |  | 145.6639
1214 | 85.825 | 232.426 |  | 147.0645
1215 | 87.585 | 234.565 |  | 147.4416
1216 | 83.48 | 229.329 |  | 146.3039
1217 | 83.501 | 218.589 |  | 135.5094
1218 | 84.418 | 229.472 |  | 145.5065
1219 | 84.196 | 230.859 |  | 147.1236
1220 | 87.665 | 231.535 |  | 144.3218
1221 | 85.705 | 230.199 |  | 144.9478
1222 | 84.72 | 229.256 |  | 144.9899
1223 | 85.834 | 229.185 |  | 143.8042
1224 | 84.208 | 228.866 |  | 145.1154
1225 | 84.623 | 228.873 |  | 144.7092
1226 | 83.062 | 220.711 |  | 138.0872
1227 | 108.698 | 243.773 |  | 135.505
1228 | 84.303 | 232.525 |  | 148.6938
1229 | 85.905 | 222.026 |  | 136.5543
1230 | 93.152 | 194.516 |  | 101.7158
1231 | 85.607 | 197.893 |  | 112.6783
1232 | 92.279 | 195.461 |  | 103.5425
1233 | 83.04 | 184.767 |  | 102.0824
1234 | 86.551 | 190.592 |  | 104.4045
1235 | 100.682 | 199.854 |  | 99.5185
1236 | 85.386 | 197.202 |  | 112.2067
1237 | 94.165 | 201.403 |  | 107.6127
1238 | 83.604 | 228.61 |  | 145.5126
1239 | 84.219 | 229.69 |  | 145.9792
1240 | 84.09 | 229.743 |  | 146.1619
1241 | 85.327 | 229.363 |  | 144.5426
1242 | 83.94 | 227.262 |  | 143.826
1243 | 84.44 | 232.127 |  | 148.2064
1244 | 85.153 | 232.045 |  | 147.5013
1245 | 85.396 | 233.45 |  | 148.6719
1246 | 86.212 | 231.013 |  | 145.4016
1247 | 87.368 | 231.58 |  | 144.8102
1248 | 86.97 | 230.702 |  | 144.3246
1249 | 86.559 | 234.463 |  | 148.5137
1250 | 85.781 | 223.518 |  | 138.3083
1251 | 85.397 | 229.158 |  | 144.3573
1252 | 84.162 | 231.422 |  | 147.8746
1253 | 87.069 | 231.939 |  | 145.4709
1254 | 87.743 | 232.532 |  | 145.3896
1255 | 85.341 | 227.868 |  | 143.1182
1256 | 85.627 | 229.883 |  | 144.8434
1257 | 87.666 | 230.877 |  | 143.7941
1258 | 86.069 | 234.63 |  | 149.1659
1259 | 86.539 | 229.128 |  | 143.1696
1260 | 86.115 | 228.536 |  | 143.0009
1261 | 86.673 | 230.709 |  | 144.6225
1262 | 86.613 | 230.479 |  | 144.4554
1263 | 85.268 | 229.687 |  | 145.0107
1264 | 87.292 | 230.166 |  | 143.463
1265 | 84.729 | 227.315 |  | 143.1738
1266 | 86.574 | 233.623 |  | 147.659
1267 | 84.452 | 231.858 |  | 148.0137
1268 | 88.196 | 232.589 |  | 144.9883
1269 | 86.309 | 236.813 |  | 151.1206
1270 | 84.792 | 228.653 |  | 144.4504
1271 | 87.423 | 230.997 |  | 144.1623
1272 | 86.461 | 230.241 |  | 144.3691
1273 | 85.985 | 229.198 |  | 143.7998
1274 | 87.495 | 231.591 |  | 144.6864
1275 | 85.753 | 228.318 |  | 143.1527
1276 | 87.072 | 231.46 |  | 144.9796
1277 | 87.745 | 231.084 |  | 143.9263
1278 | 85.887 | 233.766 |  | 148.4849
1279 | 85.94 | 229.422 |  | 144.0699
1280 | 86.437 | 229.86 |  | 143.9143
1281 | 86.567 | 233.305 |  | 147.2406
1282 | 85.406 | 229.014 |  | 144.0999
1283 | 85.42 | 228.142 |  | 143.2141
1284 | 86.484 | 233.635 |  | 147.6584
1285 | 86.142 | 234.772 |  | 149.1392
1286 | 85.277 | 228.65 |  | 143.8674
1287 | 86.618 | 233.094 |  | 146.9811
1288 | 86.598 | 225.779 |  | 139.6609
1289 | 86.217 | 229.985 |  | 144.2637
1290 | 84.921 | 227.862 |  | 143.4339
1291 | 85.499 | 229.048 |  | 144.044
1292 | 83.631 | 228.088 |  | 144.9617
1293 | 86.25 | 228.786 |  | 143.034
1294 | 84.639 | 228.715 |  | 144.5794
1295 | 85.533 | 229.109 |  | 144.0776
1296 | 86.521 | 229.536 |  | 143.5146
1297 | 86.869 | 229.623 |  | 143.2527
1298 | 85.344 | 231.807 |  | 146.9747
1299 | 84.272 | 228.643 |  | 144.8754
1300 | 86.527 | 229.987 |  | 143.9612

## Blanks
Reagent blanks
Date | R1 | R2 | Rb
2001-01-24 | 65.1 | 65.6 | -0.5
2001-01-24 | 66.6 | 66.1 | 0.5
2001-01-24 | 65.85 | 65.6 | 0.25
2001-01-24 | 61.85 | 61.6 | 0.25
2001-01-24 | 63.6 | 64.85 | -1.25
2002-04-13 | 65.4 | 65.5 | -0.1
|  | Avg | -0.14167
|  | SD | 0.58837
Avg | 64.7333 | 64.75 | -0.016667
SD | 1.5763 | 1.6248

## Standards
Iodate standard was prepared at LDEO on Oct. 20, 2000.
WKIO3 | 0.3569 |  | NIO3(tp) | 0.010008
fbuoy | 1.0002
V(tp) | 1
V(20) | 1
tp | 24
Standards based on assumption that Wheaton dispenser delivered 10.005 ml of standard.
Date | Stds | [5 empty] | Avg | SD
2000-11-22 | 650.98 | 649.67 | 649.83 |  |  |  | 650.16 | 0.5835
2000-12-09 | 653.55 | 649.78 |  |  |  |  | 651.665 | 1.885
2000-12-21 | 643.33 | 643.15 | 644.09 |  |  |  | 643.5233 | 0.40738
2000-12-23 | 657.06* | 642.51 | 642 | 642.25 |  |  | 642.2533 | 0.25502
2000-12-24 | 646.41 | 646.61 | 646.4 | 646.54 |  |  | 646.5167 | 0.10693
2000-12-26 | 643.53 | 644.76 | 645.35 | 645.24 |  |  | 645.1167 | 0.31374
2000-12-26 | 642.45 |  | 644.87 | 642.33 | 646.66 |  | 643.6 | 1.7961
2000-12-29 | 643.59 | 650.4* | 645.58 | 646.62 |  |  | 646.1 | 0.73539
2000-12-30 | 653.14* | 651.2* | 644.07 | 645.77 |  |  | 644.92 | 1.2021
2000-12-31 | 650.2 | 648.75 | 645.59 | 647.21 | 647.94 | 649.19 | 647.1833 | 1.5802
2000-12-31 | 643.85 | 643.95 | 644.81 |  |  |  | 644.2033 | 0.52776
2001-01-01 | 649.18* | 644.08 | 644.63 |  |  |  | 644.355 | 0.38891
2001-01-06 | 642.58 | 644.1 | 647.36* | 642.84 |  |  | 643.34 | 1.0748
New bottle
2001-01-09 | 645.08* | 647.08 | 648.91 | 645.75* | 648.43 | 648.12 | 647.995 | 1.294
2001-01-14 | 653.08 | 648.22 | 656.68* | 650.29 | 651.85 | 649.22 | 650.65 | 3.4365
2001-01-21 | 649.33 | 645.33 | 646.83 | 652.8 |  |  | 648.5725 | 2.8282
646.2596
LMG 0203 | [6 empty] | 2.7738
2002-04-14 | 640.71 | 646.09 | 644.35 | 648.3 |  |  | 644.8625 | 2.7762 | 0.43051
2002-04-15 | 647.27 | 647.27 | 647.33 |  |  |  | 647.29 | 0.028284
2002-04-17 | 645.04 | 644.17 | 645.67 |  |  |  | 644.96 | 0.61498 | 0.095352
2002-04-19 | 644.01 | 644.49 | 644.82 |  |  |  | 644.44 | 0.33257 | 0.051605
2002-04-22 | 645.47 | 645.88 | 646.21 |  |  |  | 645.8533 | 0.30269 | 0.046867
2002-04-24 | 645.22 | 647.56 | 645.69 |  |  |  | 646.1567 | 1.0107 | 0.15642
2002-04-27 | 645.67 | 644.37 | 644.86 |  |  |  | 644.9667 | 0.53606 | 0.083114
2002-05-02 | 646.2 | 648.54 | 645.6 |  |  |  | 646.78 | 1.2684 | 0.19611
2002-05-05 | 646.82 | 645.34 | 644.24 |  |  |  | 645.4667 | 1.0571 | 0.16377
2002-05-08 | 645.09 | 649.28 | 647.04 |  |  |  | 647.1367 | 1.7119 | 0.26454
2002-05-10 | 648.59 | 648.96 | 647.31 |  |  |  | 648.2867 | 0.70693 | 0.10905