CTD Operations and Processing

The CTD system worked well on GLOBEC III, being operated primarily by the grantees.  Marine ET Sheldon Blackman oversaw its operation on the night shift and trained the Grantees on cast procedures.  Marine ET Romeo LaRiviere oversaw operations on the day shift.  All 92 scheduled stations were completed with 95 casts.    Twelve (12) additional casts were done at areas of interest, giving a total of 107 CTD casts for Globec III.

##### SENSOR CONFIGURATION

Dual Temperature and Conductivity sensors and pumps were put on the frame, as well as a PAR sensor, a Chelsea Fluorimeter and a Transmissometer.  The PAR sensor was removed for all casts over 1000 meters and the Transmissometer was removed for all casts over 2000 meters.  (For channel connections to the Underwater Unit, See pages 3 & 4.)  (For sensor location and distance from the pressure sensor, see Page 4).

A Fast Repetition Rate Fluorimeter (FRRF) was attached to the CTD frame on casts less than 500m.

A CMiPS  (CTD-mountable Microstucture Profiling System was mounted on the CTD frame for most casts.  Temperature Sensor, T1 was connected to Voltage Channel #6 on the CTD.  (For channel connections to the Underwater Unit, See page 4.)

A new  SBE-43 Oxygen sensor was mounted on the frame.  After considerable investigation, it was discovered that the new sensor required the latest version of Seasave for Windows (Version 5_25A) in order to collect the data.  Also, the old Oxygen sensor cables would not work with the new sensor.  Therefore, a cable needed to be fabricated, as the cables that came with the sensor were too short.

A bottom contact switch and a O.L.S. pinger were added to prevent bottom contact.  The pinger started giving a very faint trace on the EPC plotter, which turned out to be due to low batteries.  It was replaced with one of the spares and the EPC trace then became much darker and easier to read.

##### BOTTLES

There were 24 bottles mounted on the CTD Rosette.  When it was used, the FRRF was mounted in place of Bottle #23.

Early on in the cruise, there was a problem with the sample bottles misfiring.  Over a number of casts the same bottles would not fire.  The problem was finally linked to the titanium trigger mechanisms.  It appears that the older mechanisms started to exhibit an inability to respond adequately to the electromagnets.   Replacing a trigger mechanism with a new or newer looking one, generally solved the problem.  After repeatedly removing the carousel and replacing trigger mechanisms with some from the spare carousel, bottle misfires were eliminated.

##### DATA ACQUISITION & PROCESSING

Both CTD computers worked well during the cruises.  Seasave for Window Ver. 5_25A was downloaded in Punta Arenas on the Web from SeaBird and installed on the acquisition computer.   The batch files on the processing computer were modified by Suzanne O’Hara to accommodate using Seasave and SeaProcess for Windows. With the new O2 sensor, it wasn’t  possible to exactly match previous cruises processing. The SeaSoft windows versions of the programs do the steps in a different order from the DOS versions, so the same order of output columns couldn’t be generated.  The new Oxygen sensor (SBE-43) also generates slightly different output columns.

Seasave for Windows turned out to be quite easy to use once the operators got used to it.  A set of  cast instructions were generated and an electronic copy is available on the  NBP Network at:

O:\et_dir\documents\ctd\Cast_WinInstructions.doc

The configuration files for the cruise are located on the acquisition computer in:

C:\seasoft\NBP0202

The soak time for the CTD was greatly reduced due to the use of the new oxygen sensor.  The sensor stabalized within 5 minutes.  However, there appears to be some time lag problems associated with the sensor that are greater than specified in the SeaBird documentation.  The difference between the upcast and downcast oxygen readings varied as much as  0.5 ml/l.  SeaBird was contacted about this problem.  Dave Armstrong at SeaBird stated that ‘The oxygen data needs to be aligned. I would suggest advancing the oxygen data by 7 to 10 seconds, this should correct the data’. Senior Analyst, Suzanne O’Hara ran some tests with time advancing of the data.  These tests resulted in the data being brought closer together, but did not align the data adequately.

To verify the oxygen sensor time lag problem, a second SBE-43 Oxygen sensor was mounted on the CTD frame and connected to ExternalVoltage Channel 2 in place of the PAR sensor for casts #070 – #073.  The second oxygen sensor exhibited the same time lag problem.    A copy of the raw data files is being forward to SeaBird for their analysis.

John Klinck, the head oceanographer on this cruise, has stated that he is going to correct the data himself using the bottles samples that he has collected. This is the way that the data for the Beckman style sensors has been corrected in the past.  Therefore, no time advances have been applied to the oxygen data in the standard processing.

Also, after all CTDs were completed, using the Term11 program, Sheldon Blackman advanced Voltage Channel 0 (O2 Sensor) on the CTD Deck Unit by 9.9 seconds and re-acquired Cast #107 from the cassette tape.  The plot from the re-acquistion was  then compared to the original acquisition.  As with the post processing tests, the downcast and upcast traces for the SBE 43 Oxygen sensor were brought significantly closer together but were still off by quite a bit.  The Term11 program only allows the deck unit to be advanced by a maximum of 9.999 seconds. The time advance on the deck unit was put back to 0 seconds after the tests were completed.

The raw data from Cast #38 had to be recovered from the audio cassette tape.  The raw data file (p202038.dat) became corrupt when an attempt was made to edit the header in the file on the acquisition computer.  The file was transferred over to the network and subsequently to the processing computer.  An attempt to process the file revealed its corrupt status.  Fortunately, the file was recovered from the backup audio tape.  All processing after that went well with the exception of the sound velocity file (sp20238.cnv) file did not process properly.  For some reason there is still a problem with raw data files that prevents this file from being created properly.

There were occasional problems with Seasave starting the acquisition if the NMEA GPS string was not there.  Occasionally the Ashtech GPS would fail to be  locked on a GPS fix and would not be sending a NMEA string.  Resetting the Ashtech GPS system solved the problem and the NMEA string was then available.  However, the Seasave software sometimes needed to be rebooted when this occurred or the software would still fail to start the acquisition.   No data was lost due to this problem.  By making sure the Ashtech was operating immediately before the cast started, the problem was minimized.

# SBE-9 Plus Underwater Unit

#### S/N  09P10716-0377

Primary Conductivity

S/N 2067

## Primary

## Temperature

S/N 2367

## Secondary

## Temperature

S/N 2186

Secondary Conductivity

S/N 2513

### Digiquartz

Pressure Sensor

S/N 58949

Bottom End Cap Connector

(Configuration on NBP0202)

# SBE-9 Plus Underwater Unit

S/N  09P10716-0377

Ext. Volt 6: CMiPS__

Ext. Volt 7: ________

Ext. Volt 4: Transmissometer, S/N 207D

Ext. Volt 5: Flourimeter, S/N 080880

Ext. Volt 0: Oxygen SBE43,S/N 0080

Ext. Volt 1: _____________________

Ext. Volt 2: PAR/Irradiance, S/N 4361

Ext. Volt 2: Oxygen SBE43,S/N 0082*

Ext. Volt 3: _____________________

*  Note: Secondary Oxygen sensor replaced

PAR on casts #070 – 073.

# Top End Cap Connector

(Configuration on NBP0202)

CONFIGURATION for NBP0202 CTD

ASCII file: C:\seasoft\NBP0202\Nbp0202.con or Nbp0202_Ox.con (with secondary Oxygen sensor)

Configuration report for SBE 911/917 plus CTD

---------------------------------------------

Frequency channels suppressed : 0

Voltage words suppressed      : 0

Computer interface            : RS-232C

Scans to average              : 1

Surface PAR voltage added     : No

NMEA position data added      : Yes

Scan time added               : No

1) Frequency channel 0, Temperature

Serial number : 2367

Calibrated on : 12-JUN-01

G             : 4.33466091e-003

H             : 6.43057120e-004

I             : 2.37105286e-005

J             : 2.29326732e-006

F0            : 1000.000

Slope         : 1.00000000

Offset        : 0.0000

2) Frequency channel 1, Conductivity

Serial number : 2067

Calibrated on : 13-DEC-01

G             : -1.03980803e+001

H             : 1.46700966e+000

I             : -5.12766449e-003

J             : 4.41800486e-004

CTcor         : 3.2500e-006

CPcor         : -9.57000000e-008

Slope         : 1.00000000

Offset        : 0.00000

3) Frequency channel 2, Pressure, Digiquartz with TC

Serial number : 58949

Calibrated on : 13-JUL-01

C1            : -4.840395e+004

C2            : -2.017057e-003

C3            : 1.464810e-002

D1            : 3.990600e-002

D2            : 0.000000e+000

T1            : 2.998386e+001

T2            : -2.560542e-004

T3            : 3.869120e-006

T4            : 2.452640e-009

T5            : 0.000000e+000

Slope         : 0.99999000

Offset        : 0.21880

AD590M        : 1.146000e-002

AD590B        : -8.457340e+000

4) Frequency channel 3, Temperature, 2

Serial number : 2186

Calibrated on : 15-DEC-01

G             : 4.34039770e-003

H             : 6.45080125e-004

I             : 2.35267825e-005

J             : 2.25318679e-006

F0            : 1000.000

Slope         : 1.00000000

Offset        : 0.0000

5) Frequency channel 4, Conductivity, 2

Serial number : 2513

Calibrated on : 13-DEC-01

G             : -1.05905818e+001

H             : 1.63509839e+000

I             : -2.18814632e-003

J             : 2.63682166e-004

CTcor         : 3.2500e-006

CPcor         : -9.57000000e-008

Slope         : 1.00000000

Offset        : 0.00000

6) Voltage channel 0, Oxygen, SBE 43

Serial number : 0080

Calibrated on : 02-JAN-02

Soc           : 3.8100e-001

Boc           : 0.0153

Offset        : -0.6242

Tcor          : 0.0005

Pcor          : 1.35e-004

Tau           : 0.0

7) Voltage channel 1, Free

8) Voltage channel 2, PAR/Irradiance, Biospherical/Licor

Serial number        : 4361

Calibrated on        : 12-DEC-2001

M                    : 1.00000000

B                    : 0.00000000

Calibration constant : 19801980198.01979800

Multiplier           : 1.00000000

Offset               : 0.00000000

(Replaced PAR sensor for Casts 70, 71, 72 & 73)

8) Voltage channel 2, Oxygen, SBE 43, 2

Serial number : 0082

Calibrated on : 12-JAN-02

Soc           : 3.7100e-001

Boc           : 0.0214

Offset        : -0.6270

Tcor          : 0.0023

Pcor          : 1.35e-004

Tau           : 0.0

9) Voltage channel 3, Free

10) Voltage channel 4, Transmissometer, Chelsea/Seatech/Wetlab CStar

Serial number : 207D

Calibrated on : 21-JUN-2001

M             : 20.0000

B             : -0.0200

Path length   : 0.250

11) Voltage channel 5, Fluorometer, Chelsea Aqua 3

Serial number : 088080

Calibrated on : 01-FEB-2002

VB            : 0.027300

V1            : 1.860200

Vacetone      : 0.117900

Scale factor  : 1.000000

Slope         : 1.000000

Offset        : 0.000000

12) Voltage channel 6, CMiPS Temperature Sensor

13) Voltage channel 7, Free